Smart-Trak Command Set and RS232 Protocol

Important Customer Notice

This document is for firmware version 1.03 and higher.  Sierra Instruments, Inc. reserves the right to change the command set without notification to the user.  However, Sierra will also make every effort to maintain backwards compatibility with previously published command sets for Smart-Trak.  There is no implied warranty or guarantee regarding the use of this commands set.  Sierra Instruments, Inc. is not liable for any damage or personal injury, whatsoever, resulting from the use of this command set.

Protocol

The Smart-Trak communication is based on a standard RS232 port.  All bytes are ASCII except the two CRC bytes used for a redundancy check.   Each command starts with four bytes of command type, followed by a variable length value section.  The two CRC bytes follow the value string and a carriage return ends the string.  Each command string will vary in length but, can not exceed 26 bytes.

Command String = {4 command bytes + variable length value + 2 CRC bytes + carriage return} < 26 bytes.

Note: Some commands that request information to be sent back start with a “?”,  then three or four command bytes.

The RS232 port does not depend on hardware handshaking and uses only three wires on the port: transmit, receive and ground.  The port on the host needs to be configured to (9600,n,8,1) : 9600 baud, no parity,  eight bit characters, one stop bit.

Smart-Trak meters always transmit their flow data with out the host asking for a data point.  If the meter is set in analog set point mode, the flow data and the current analog set point are sent automatically.  Flow and set point data are sent about twice a second.  There is no operation mode that turns off this feature of automatic data sending.  Smart-Trak will send back the command sent to it or a group of commands that indicate new values that the sent command modified.

Quote marks indicate an ASCII string in this document.

Command descriptions:

Command headers are case sensitive.  They start with a capitol letter or with a ‘?’.

Values are signed  floating point numbers or  positive integer numbers.

CC are two cycle redundancy check bytes. Values are binary not ASCII

R stands for carriage return ( ASCII value 13 decimal)

FlowValueCCR   		Example ‘Flow3.111ccr’

Value contains the mass flow reading

Flow = 4 byte command header

Value=0 to full scale, floating point number

CC= 2 byte redundancy check

R=carriage return

Flow is a receive only command.  Smart-Trak continuously transmits this value.

SinvValueCCR		Example ‘Sinv1.456ccr’

Value contains the set point input value.  Only used if unit is a controller.

Sinv= 4 byte command header; Signal input or set point

Value=0 to full scale plus over range of meter, floating point number

CC=2 byte redundancy check

R=carriage return

If the set point is driven by an analog external source then this command is transmitted continuously.  If the source is the host computer or the Pilot Module then use this command to set the set point.

This command returns the following meter commands:

‘Unti’

‘Fscl’

‘Sinv’

‘Gasi’

‘Vlvi’

SiniValueCCR		Example ‘Sini3ccr’

Sini= 4 byte command header; Selects the source for the set point.

Value= 1 to 5 integer.

1= Module/RS232 as the set point source.

2=0-5 volt analog set point source

3=0-10 volt analog set point source

4=1-5 volt analog set point source

5=4-20 ma analog set point source

CC=2 byte redundancy check

R=carriage return

This command returns the following meter commands:

‘Unti’

‘Fscl’

‘Sinv’

‘Gasi’

‘Vlvi’

GasiValueCCR		Example ‘Gasi2ccr’

Gasi=4 byte command header; Select gas

Value=1 through 10, integer.  The following list is the factory default list. If other gases have replaced the default list, value will correspond to their order in the gas list.

1=Air

2=Argon

3=Carbon Dioxide

4=Carbon Monoxide

5=Helium

6=Hydrogen

7=Methane

8=Nitrogen

9=Nitrous Oxide

10=Oxygen

CC=2 byte redundancy code

R=carriage return

This command returns the following meter commands:

‘Unti’

‘Fscl’

‘Sinv’

‘Gasi’

‘Vlvi’

FsclValueCCR		Example ‘Fscl10.00ccr’

Fscl=4 byte command header; Set user full scale.

Value=50% of Factory full scale through 100% Factory full scale, floating point.

Factory full scale is the flow indicated in the certification sheet for the flow device.

CC=2 bytes redundancy check

R=carriage return

This command returns the following meter commands:

‘Unti’

‘Fscl’

‘Sinv’

‘Gasi’

‘Vlvi’

UntiValueCCR		Example ‘Unti2ccr’

Unti=4 byte command header; Select flow meter units.

Value=1 through 30, integer

1=scc/s

2=scc/m

3=scc/H

4=Ncc/s

5=Ncc/m

6=Ncc/H

7=SCF/s

8=SCF/m

9=SCF/H

10=NM3/s

11=NM3/m

12=NM3/H

13=SM3/s

14=SM3/m

15=SM3/H

16=sl/s

17=sl/m

18=sl/H

19=NL/s

20=NL/m

21=NL/H

22=g/s

23=g/m

24=g/H

25=kg/s

26=kg/m

27=kg/H

28=lb/s

29=lb/m

30=lb/H

CC=2 byte redundancy check

R=carriage return

This command returns the following meter commands:

‘Unti’

‘Fscl’

‘Sinv’

‘Gasi’

‘Vlvi’

OutiValueCCR		Example ‘Outi1ccr’

Outi=4 byte header.  Select output mode

Value=1 through 3, integer.

1=0-5 volt and 4-20 ma outputs

2=1-5 volt and 4-20 ma outputs

3=0-10 volt and 4-20 ma outputs

CC=2 byte redundancy check

R=carriage return

This command returns the following meter commands:

‘Unti’

‘Fscl’

‘Sinv’

‘Gasi’

‘Vlvi’

VlviValueCCR		Example ‘Vlvi2ccr’

Vlvi=4 byte header; Set valve’s operating mode.  Valid for controllers only.

Value=1 through 3, integer.

1=Automatic ,  controls to the set point value

2=Closed

3=Purge

CC=2 byte redundancy check

R=carriage return

Command is echoed back

XwrdValueCCR		Example ‘Xwrd1234ccr’

Xwrd=4 byte header;  value is 4 ASCII number characters used to set the password

Value=0000 through 9999

CC=2 byte redundancy check

R=carriage return

The flow meter stores the password, but it is up to the user application to implement a password protection scheme.   All software commands work without password protection.

The password command is echoed back.

To read a value from the flow meter without writing a value to it, place a ‘?’ in front of the four letter command followed with the CCR bytes.

‘?SinvCCR’ will return ‘Sinv’+value+CC+R;  ‘Sinvxx.xxxccr’

‘?SiniCCR’ will return ‘Sini’+x+CC+R; x is same as the Sini command

‘?XwrdCCR’ will return ‘Xwrd’+xxxx+CC+R; xxxx is the four number password

‘?VlviCCR’  will return ‘Vlvi’+x+CC+R;  x is same as the set Vlvi command

‘?OutiCCR’ will return ‘Outi’+x+CC+R; x is same as the set Outi command

‘?UntiCCR’ will return ‘Unti’+xx+CC+R; xx is same as the set Unti command

‘?FsclCCR’ will return  ‘Fscl’+value+CC+R; value is same as the Fscl command

‘?GasiCCR’ will return ‘Gasi’+xx+CC+R; xx is same as the Gasi command

The following commands are read only commands.

‘?VerCCR’ will return ‘?Ver’+string+CC+R; string = version number – may include letters

‘?FfsCCR’ will return ‘?Ffs’+value+CC+R; value = factor full scale flow, floating point number

‘?SrnCCR’ will return ‘?Srn’+string+CC+R; string= serial number – may include letters

Calculating the CC bytes.

The following is a Visual Basic version of calculating the cc bytes.

The string to calculate the cc is called crcString.  The results will be in crcWord.

EndOfString = len(crcString)

crcWord = &HFFFF   			‘ if VB thinks this is –1 then use (&H1FFFE/2)

For j = 1 To endOfString		‘ step through string

crcByte = Asc(Left(crcString, 1))	‘ get next byte

crcString = Mid(crcString, 2, endOfString - j) ' drop first char of string

crcWord = crcWord Xor (CLng(crcByte) * &H100)	‘ shift left crcByte by 1 byte and xor to

‘ crcWord

For i = 1 To 8				‘ step through the 8 high bits

If (crcWord And &H8000) Then		‘ test sign bit

crcWord = (crcWord * 2) Xor &H1021 '  shift left 1 bit then xor prime number

Else

crcWord = crcWord * 2		‘ shift left 1 bit

End If

crcWord = (((crcWord * 2) And &HFFFF	‘ mask to word length

Next i

Next j

If (crcWord And &HFF00) = &HD00 Then   	 ' carriage char not allowed

crcWord = crcWord + &H100          		 ' make this byte &h0exx

End If

If (crcWord And &HFF00) = &H0 Then    	 ' zero not allowed

crcWord = crcWord + &H100           		' make this byte &h01xx

End If

If (crcWord And &HFF) = &HD Then        	 ' carriage char not allowed

crcWord = crcWord + 1               		' make this byte &hxx0e

End If

If (crcWord And &HFF) = &H0 Then       	 ' zero not allowed

crcWord = crcWord + 1              		 ' make this byte &hxx01

End If

Visual Basic parses on carriage return and the flow meter parsers on zero making either of these byte patterns illegal to be in the cc bytes.  When creating a command to send to the flow meter, pack the high byte of crcWord first and then the low byte.

The following is a C version of calculating the cc bytes.

crc=0xffff;

while(!(crcString [j]==0){		// loop through string and calc a crc until null char

crc = crc ^ ((INT16U) (crcString [j] * 0x0100);

for(i=0;i<8;i++){

if(crc & 0x8000)

crc = (crc<<1)^0x1021;	// prime number

else

crc = crc<<1;

}

j++;

}

if((crc & 0xff00)==0x0d00) crc +=0x0100;	// increment if byte is carriage return

if((crc & 0x00ff)==0x000d) crc +=0x0001;	// increment if byte is carriage return

if((crc & 0xff00)==0x0000) crc +=0x0100;	// no zero bytes allowed

if((crc & 0x00ff)==0x0000) crc +=0x0001;	// no zero bytes allowed